# Chromatography and LCMS

Standard LCMS measurements were performed on a Dionex Ultimate3000 (Thermo Scientific, Darmstadt, Germany) using an EC100/2 Nucleoshell C18 2.7μm column (Macherey-Nagel, Düren, Germany) HPLC coupled to a micrOTOFQ II (Bruker Daltonics, Bremen, Germany) ESI-qTOF-HRMS. The LC system was run in the following gradient (A:H2O; B: MeOH, ; Flow: 200 μL/min): 0 min: 90% A, 5 min: 90% A, 35min: 0% A, 50min: 0% A, 51min: 90% A, 60 min: 90% A. The injection volume was 5 µL.

Alternatively, a setup consisting of an Agilent Infinity 1290 UPLC system (Agilent, Santa Clara, CA, USA) equipped with an Acquity UPLC BEH C18 1.7 μm (2.1 × 100 mm) column and an Acquity UPLC BEH C18 1.7 μm VanGuard Pre-Column (2.1 × 5 mm; both columns purchased from Waters, Eschborn, Germany) coupled to a DAD detector and a micrOTOFQ II mass spectrometer (Bruker Daltonics, Bremen, Germany) with an electrospray ionization source was employed. The LC system was operated using a gradient (A: H2O, 0.1% formic acid; B: MeCN, 0.1% formic acid; flow: 600 μL/min): 0 min: 95% A; 0.80 min: 95% A; 18.70 min: 4.75% A; 18.80 min: 0% A; 23.00 min: 0% A; 23.10 min: 95% A; 25.00 min: 95% A and the column oven temperature was set to 45 °C. The injection volume was 5 µL.

High accuracy measurements were performed on a 1290 UHPLC system (Agilent, Santa Clara, CA, USA) equipped with DAD, ELSD and maXis II™ (Bruker, Billerica, MA, USA) ESI-qTOF-UHRMS with the following gradient: 0 min: 95% A; 0.30 min: 95% A; 18.00 min: 4.75% A; 18.10 min: 0% A; 22.50 min: 0% A; 22.60 min: 95% A; 25.00 min: 95% A (A: H2O, 0.1% formic acid; B: acetonitrile, 0.1% formic acid; flow: 600 μL/min). The employed column was an Acquity UPLC BEH C18 1.7 μm (2.1 × 100 mm) column with an Acquity UPLC BEH C18 1.7 μm VanGuard Pre-Column (2.1 × 5 mm). The column oven temperature was maintained at 45°C. The injection volume was either 1 or 2 µL.