MS Data acquisition

Data for untargeted metabolomics using all three chromatographic methods were acquired for mass ranges of 70-1000 m/z by full MS at 70,000 resolution in both positive and negative ionization modes. The automatic gain control (AGC) and maximum injection time (IT) were 5 x 105 and 100 ms, respectively. The heated electrospray ionization (HESI) source was operated with sheath gas, 15 arbitrary units; auxiliary gas, 5 arbitrary units; capillary temperature, 250oC; auxiliary gas heater temperature, 50oC; and S-lens RF level, 50. The spray voltage was 4.2 and 3.9 kV in positive and negative modes, respectively. One sample in each group was also used for the top 12 data-dependent acquisition experiments in both ionization modes to generate MS/MS datasets for compound identification. These experiments involved a full MS scan at 70,000 resolution, AGC target 5 x 105, maximum IT 100 ms and MS/MS scans at 17,500 resolution, AGC target 5 x 104, maximum IT 50, 2.0 m/z isolation window, stepped collision energy of 15, 25 and 35 eV, intensity threshold 1 x 104 and 15-sec dynamic exclusion.